VMAT TRF(Treatment Delivery System) file analysis with RTplan DICOM(TPS)

- Target Dose: 61.6Gy/28fx

- Total MU: 643.5

- Number of control points: 147

- Max. Gantry Speed (TPS): 6 deg/s

- Max. Dose Rate (720MU/min)

- Max. num of Arc: 1

Inspired by Park, J. M., they consider leaf speed, leaf acceleration, gantry speed and dose rate in QA metrics, found the MI_s, MI_a, MI_t had high correlation with gamma passing criteria results in VMAT plan. But the conventional QA metrics like MCS/LTMCS seems in-correlated with VMAT QA results.

(Mateirals):

- 20 NPC and 20 Prostate VMAT plans

- Eclipse System (two full arcs), dose grid 2.5mm

- MapCHECK2 (measured 2D dose distribution V.S. TPS calculated dose)

We aim to use RTPlan DICOM to estimate the gantry speed, MLC speed, acceleration and dose rate during the delivery. But the difference between PLAN DICOM and TRF indicates the estimate for MLC speed, gantry speed is far away from the reality.

- Check the consistency of differential MU between TPS and TDS

- Check time interval between control points (estimated from DICOM VS. Delivery).

- Check average gantry speed between control points (estimated from time interval). Though gantry speed exists negative, the dose rate is 0 in this gantry position. So it won’t influence the final delivery dose.

- Check average dose rate between control points (estimated from time interval)

- Check MLC speed between 2 control points (from TRF e.g. CP3-> CP4)

Reference

[1] Park, J. M., Park, S. Y., Kim, H., Kim, J. H., Carlson, J., & Ye, S. J. (2014). Modulation indices for volumetric modulated arc therapy. Physics in Medicine & Biology, 59(23), 7315.