VMAT TRF(Treatment Delivery System) file analysis with RTplan DICOM(TPS)

- Target Dose: 61.6Gy/28fx

- Total MU: 643.5

- Number of control points: 147

- Max. Gantry Speed (TPS): 6 deg/s

- Max. Dose Rate (720MU/min)

- Max. num of Arc: 1

Inspired by Park, J. M., they consider leaf speed, leaf acceleration, gantry speed and dose rate in QA metrics, found the MI_s, MI_a, MI_t had high correlation with gamma passing criteria results in VMAT plan. But the conventional QA metrics like MCS/LTMCS seems in-correlated with VMAT QA results.

(Mateirals):

- 20 NPC and 20 Prostate VMAT plans

- Eclipse System (two full arcs), dose grid 2.5mm

- MapCHECK2 (measured 2D dose distribution V.S. TPS calculated dose)

- Check the consistency of differential MU between TPS and TDS

- Check time interval between control points (estimated from DICOM VS. Delivery)

- Check average gantry speed between control points (estimated from time interval)

- Check MLC speed between 2 control points (from TRF e.g. CP3-> CP4)

Reference

[1] Park, J. M., Park, S. Y., Kim, H., Kim, J. H., Carlson, J., & Ye, S. J. (2014). Modulation indices for volumetric modulated arc therapy. Physics in Medicine & Biology, 59(23), 7315.